| ## Member State of OIML Germany |  | OIML Certificate No. R76/1992-DE1-05.10 Revision 2 |
| --- | --- | --- |

| OIML CERTIFICATE OF CONFORMITY |  |
| --- | --- |
| Issuing Authority |  |
| Name: | Physikalisch-Technische Bundesanstalt |
| Address: | Bundesallee 100 38116 Braunschweig |
| Person responsible: | Dr. Oliver Mack |

Applicant

| Name: | seca gmbh & co. kg |
| --- | --- |

| Address: | Hammer Steindamm 9 - 25 22089 Hamburg Germany |
| --- | --- |

Manufacturer of the certified type is the applicant.

| Identification of the certified type | Non-automatic electromechanical weighing instrument for persons Types: M704x2, M764x2, COS01A, COS01B, COS01C Further characteristics see page 2 |
| --- | --- |
| This Certificate attests the conformity of the above identified type (represented by the sample or samples identified in the associated Test Report) with the requirements of the following Recommendation of the International Organization of Legal Metrology (OIML): |  |
| R76-1, edition 1992, including Amendment 1 (1994), for accuracy class |  |
| This Certificate relates only to the metrological and technical characteristics of the type of instrument covered by the relevant OIML Recommendation identified above. This Certificate does not bestow any form of legal international approval. |  |

This Revision 2 has been issued because new variants of types COS01B and COS01C with new electronic has been added. The conformity is established in the Report No. 1.12-4017690, Revision 2 (11 pages). The metrology results of tests and examinations provided in the Test Reports No. 1.12-4017690/2 (51 pages) and 1.12-4017690/3 (38 pages).

The test results of the former Test Reports remains valid.

| The Issuing Authority | The CIML Member |
| --- | --- |
| Dr. O. Mack | Dr. R. Schwartz |
| Head of Working Group | Head of Division |

| 31.10.2012 | 31.10.2012 |
| --- | --- |

L.S.

Identification of the pattern (continued)

The weighing instrument consists of a weighing platform with one strain gauge shear-beam load cell, an indication for displaying the weighing result and a membrane keyboard to operate the instrument.

The weighing ranges with Max, Min, e, d and number of verification scale intervals may be chosen within the limits of No. 3.2 of R 76-1 and of the table 1.

## Table 1

| Type | M704x2 | M764x2 | COS01A |
| --- | --- | --- | --- |
| Accuracy class |  |  |  |
| Min | 2 kg | 2 kg | 2 kg |
| Max | 250 kg | 250 kg | 250 kg |
| e=d | 0,1 kg | 0,1 kg | 0,1 kg |
| n | 2500 | 2500 | 2500 |
| Tare balancing range, subtractive |  100 % of Max |  | -- |
| Temperature range | 10 °C ... 40 °C |  |  |
| Platform medium sized | 308 mm x 345 mm |  |  |

## Table 1 (continue)

| Type | COS01B | COS01C |
| --- | --- | --- |
|  | Multiple range instrument |  |
| Accuracy class |  |  |
| Max (Max1 ¦ Max2) | 150 kg ¦ 300 kg | 150 kg ¦ 250 kg |
| e=d= (e1 ¦ e2) | 0,05 kg ¦ 0,1 kg | 0,05 kg ¦ 0,1 kg |
| n= (n1 ¦ n2) | 3000 ¦ 3000 | 3000 ¦ 2500 |
| Subtractive tare balancing device | -- |  |
| Temperature range | 10°C … 40°C |  |
| Weighing platform (on average) | 308 mm x 345 mm |  |

Important note: 	Apart from the mention of the Certificate's reference number and the name of the OIML Member State in which the Certificate is issued, partial quotation of the Certificate and of the associated Test Report(s) is not permitted, although either may be reproduced in full.