SUMMARY OF COMMENTS

TO THE 2 CD OF IR “TRANSMITTERS WITH ELASTIC SENSING ELEMENTS”

|  | Points | Comments | Country | Answers |
| --- | --- | --- | --- | --- |
| 1 | - TERMINOLOGY | Definition in the document “Gauge pressure: the pressure greater than ambient pressure, the latter being considered as the datum point.” Hence, the gauge pressure is the difference between the absolute pressure and the given (actual absolute) atmospheric pressure. This difference can be positive or negative | Austria | Accepted |
| 2 |  | The term “Vacuum gauge pressure” – is not a common term or special type of pressure | Austria | Accepted |
| 3 |  | Verification: We would suggest to use the wording of VIML, Edition 2000) | Austria | Accepted |
| 4 |  | Maximum permissible error 3.5 We would prefer the definition of VIM, JCGM 200:2008 | Austria | Accepted |
| 5 |  | - Gauge pressure Instead of: A pressure greater than ambient pressure, better: A pressure greater or less than ambient pressure " (cf. 2.3 of EN 472 "Pressure gauges – May be better: - Gauge pressure The difference between absolute pressure pabs and the given (absolute) atmospheric pressure pamb is the atmospheric pressure difference pe; is termed gauge pressure pe = pabs − pamb - Gauge pressure pe assumes positive values when the absolute pressure is greater than the atmospheric pressure; it assumes negative values when the absolute pressure is less than the atmospheric pressure. NOTE In practice, the term gauge pressure or “relative pressure” is also used for pressure measuring equipment in the following cases, where the reference pressure may be: a) The given (absolute) atmospheric pressure pamb b) The (given) atmospheric pressure pamb present at the time of manufacture of the pressure measuring equipment c) A fixed pressure value, which usually corresponds to an absolute pressure pabs ≈ 1 bar Case a) corresponds to gauge pressure. In cases b) and c), the reference source is not open to the respective (absolute) atmospheric pressure; it is a closed volume (sealed gauges). Absolute pressure Absolute pressure pabs is the pressure compared to the zero pressure in a molecular-free volume Differential pressure Difference between two pressures p1 and p2, when it is itself a measurand: Δp = p1 – p2 pressure difference or also differential pressure p1,2 | Germany | Accepted |
| 6 |  | Suggest the term “sensivity” is used without “permission”. Also align with VIM 4.12 | Unites Kingdom | Accepted |
| 7 |  | Vacuum gauge pressure The pressure less than ambient pressure, the latter being considered as the datum point Sometimes Vacuum gauge pressure is used for absolute vacuum gauges. Maybe it should be necessary to define “Absolute vacuum gauge”. | Spain | Accepted |
| 8 |  | Vacuum gauge pressure is not a type (kind) of pressure. | Germany |  |
| 9 |  | The electric output signal that is the result of the electric signal conversion coming from primary pressure transmitter into direct current standard signal of current: (0 - 5), (0 - 20), (4 - 20) mA; of voltage: (0 - 1), (0 - 5) V and others, if they are allowed by National Regulations. Put “Transducer” instead of “transmitter”. | Spain | Not accepted. The term “transmitter” is a part of the term “transducer” (see, for example, terminology in SAMA STANDARD PMC20-2-1970) |
| 10 |  | Calibration Operation, that, under specified conditions, in a first step establishes a relation between the quantity values with measurement uncertainties provided by measurement standards and corresponding indications with associated measurement uncertainties and, in a second step uses this information to establish a relation for obtaining a measurement result from an indication. It is defined in VIM. I thing, It’s necessary to define here. | Spain | Not accepted This definition aligns with point 2.39 of VIM 2010 |
| 11 |  | Pin - input value (for example, pressure) at transmitter input Are there any case, when the input value it isn’t pressure? | Spain | For example, empiric imitation of signal, received from thermopile |
| 12 |  | Measurement uncertainty It is defined in VIM. I thing, It’s necessary to define here. | Spain | Not accepted This definition aligns with point 2.26 of VIM 2010 |
| 13 |  | p. 3 unclosed parenthesis | Czech Republic | Accepted |
| 14 |  | Suggest deleting “concrete” and inserting “particular”. The same comment applies at 4.1.4, 4.1.1.0, and 4.2.2.2. Correct “from” to “form” | United Kingdom | Accepted |
| 15 | p. 2.1 | a) Can the last sentence be clarified please? | United Kingdom | The transmitters should be equipped with indicating, digital or other measurement devices including non contact (Bluetooth, WiFi, etc). Maybe it is inexactly formulized in English. We mean the presence of additional output devices except the basic one: 4…20, 0…20, 0…5 (mA). The output signal in the form of Fieldbus foundation should be also added here. |
| 16 |  | b) Suggest for second hyphen something like: “Channels to measure ambient pressure or temperature.” Suggest for third hyphen something like: “…In the form of a computing device whose main calculating ….” | United Kingdom |  |
| 17 |  | 2. DESCRIPTION OF TRANSMITTER a) Instead of: transmitters for gauge (including hydrostatic)* and absolute pressure**……… Without: transmitters for gauge (including hydrostatic)* and absolute pressure …. (without footnote) | Germany | Accepted |
| 18 |  | Is the additional equipped part of the transmitter? Maybe, we should say that “The additional equipments are out of the pattern evaluation”. | Spain | Not accepted The additional equipment is a part of the transmitter and it may be examined together with the basic equipment according to the test program. |
| 19 |  | p.6 one more parenthesis | Czech Republic | Accepted |
| 20 | p. 3.2 | Delete “(for example kgf/cm2)” - kgf/cm2 is no SI unit | Austria | Not accepted. The unit kgf/cm2, mm Hg and other units may be used in accordance with National Regulations We would like to keep kgf/cm2 on because this unit is widely used in Russia. |
| 21 |  | 3. THE UNITS OF PRESSURE MEASUREMENT instead of: The bar and its submultiples, especially the mbar, as well as other units (for example kgf/cm2), may be used insofar without: (for example kgf/cm2) kgf/cm² is not SI unit and not legal unit in Europe and (I think) USA | Germany |  |
| 22 |  | Why can the kgf/cm² used? Maybe We can use mm Hg for medical application but: What is the use allowed to “kgf/cm”? | Spain |  |
| 23 | p. 4.1.3 | Suggest this includes the standard European supply of 230 VAC. | United Kingdom | Accepted |
| 24 | p. 4.1.6 | Suggest something like: “The transmitter shall be capable of indicating measured values either directly or remotely. Indications may be analogue or digital.” | United Kingdom | Accepted |
| 25 |  | Why? We think. One thing is the transducer and other is the indicators. Both of them are independents. | Spain | Not accepted Additionally to electrical output signal some transmitters are provided by other analogue or digital output devices, agreed with the customer. |
| 26 | p. 4.1.11 | electrical schematic of the pressure transmitter. | Spain | Not accepted The construction of transmitters in most cases don’t let place the electrical schematic on the transmitter’s body. In this instance it has to be placed on the operating guide. |
| 27 | p. 4.2.1.1 | What do transmitter with elastics sensor have this accuracy of ± 0,02? | Spain | Not accepted Nowadays the transmitters with reducial error ± 0,025% (for example, transmitters, which are produced by Emerson company). |
| 28 | p. 4.2.2.2. | instead of: -atmospheric (barometric) pressure for transmitters of absolute pressure, without: for transmitters of absolute pressure, or better: -atmospheric (barometric) pressure for transmitters of absolute pressure for use as gauge pressure transmitter, | Germany | Not accepted Some transmitters of absolute pressure (especially it refers to sealed-gauge transmitters) are exposed to the influence of atmospheric pressure. |
| 29 | p. 4.2.2.5 | We think that in the formula given here a special term for the influence mounting is missing, especially the influence of the torque, with which the transducer is screwed. | Austria | Accepted |
| 30 |  | better: permit also additional factors of influence - Ywc=Kp⋅γt2+γu2+γr2+γd2+γn2+γx2, where - γt – change of output signal of the transmitter due to change of ambient air temperature. - γu – change of output signal of the transmitter due to voltage change. - γr – change of output signal of the transmitter due to readjustment. - γd – change of output signal of the transmitter due to static pressure change. - γn – change of output signal of the transmitter due to vibration. - γx – change of output signal of the transmitter due to other influences (for example position change influence or other) - Kp – coverage factor depended on reliability requirements. | Germany | Accepted |
| 31 | p. 4.4 | Suggest the term “sensitivity” is used without “permission”. Also align with VIM 4.12. | United Kingdom | Accepted |
| 32 | p. 5.2 | „hazardous“ instead „hazard“ | Czech Republic | accepted |
| 33 | p. 7.1 | Must every manufacture do a pattern evaluation or approval of every range, every accuracy, every kind of record. Etc…? We think. The manufactures can not do the evaluation of all this transducers. There area too much combinations. It is economically impossible. What is the point of view to select the number of the different kinds of transmitters? | Spain | Not accepted The tests should be conducted on the typical transmitters with maximum accuracy, having minimal and medium ranges. You shouldn’t conduct metrological tests on transmitters, which differ, for example, from forms or the material of the body. |
| 34 | p. 7.9.1 | 65% ± 5 % is too much high | Spain | Not accepted This point aligns with point 5.3 of IEC 770 |
| 35 |  | Why is upper limit 106.7 kPa? Such accuracy is nonsense. 107 kPa is enough (if there even some testing laboratory at shores of Dead Sea). Also recommended 101.3 kPa is nonsense. Not all labs are at sea level. | Czech Republic | Not accepted. Recommended by IEC 770 range of atmospheric pressure is 860 mbar (86 kPa) to 1060 mbar (106 kPa). |
| 36 | p. 7.10.1 | We think is enough with 6 point including zero and maximum pressure value for all kind of transmitter. | Spain | Not accepted We agree that it is possible, but the majority of OIML members consider that for the high-accuracy transmitters it’s necessary to enlarge the quantity of points in order to decrease the interpolation errors. |
| 37 | 7.10.4 | h. The check of the atmospheric pressure variation influence to absolute pressure transmitters The check of the atmospheric pressure variation influence to the output signal of absolute pressure transmitters is carried out in the barometric chamber with one fixed value of measured pressure and with pressure of ambient air in the pressure chamber equal to atmospheric. Changing chamber pressure for 10 kPa the output signal variation is determined. The transmitter is considered as having passed the test successfully, if its characteristics don’t exceed maximum permissible values, specified in the Technical (and advertising) Documents of the Manufacturer. instead of: h. The check of the atmospheric pressure variation influence to absolute pressure transmitters without: h. The check of the atmospheric pressure variation influence to absolute pressure transmitters or better: h. The check of the atmospheric pressure variation influence to absolute pressure transmitters for use as gauge pressure transmitter - This test may be for testing what type of pressure is the transmitter. - For example for testing: Is the transmitter a gauge pressure transmitter or an absolute pressure (for example “sealed gauge”) transmitter. - For absolute pressure transmitter, there is no influence of atmospheric pressure! - This test must be for absolute pressure transmitter for use as gauge pressure transmitter. | Germany | Not accepted Some transmitters of absolute pressure (especially it refers to sealed-gauge transmitters) are exposed to influence of atmospheric pressure. |
| 38 |  | g. The check of the static pressure effects The check of the working pressure variation influence is carried out for differential pressure transmitters. The check is separately carried out for zero point and for other points (for the inclination angle) of actual transfer function Remove “(for the inclination angle)” | Spain | Accepted |
| 39 |  | i. The check of the ripple content of electrical output of the transmitter with current output signal The check of the ripple content of electrical output of the transmitter with current output signal is carried out with an oscilloscope by the transmitter voltage alternating component at the upper and lower ends of the span and with maximum load resistance 250 Om and 1000 Om for transmitter with upper values of output signal 20 mA and 5 mA Change “Om” by “Ohm”. | Spain | Accepted |
| 40 |  | - The check of the increased relative humidity influence corresponding to transport conditions. The check of the increased humidity influence to the packed transmitters according to the requirements of the Technical Documents is carried out as follows. The transmitters are placed into a heat and humidity chamber and relative humidity increases up to 95 % at 35 C Too much high value. We propose 90 % | Spain | Accepted |
| 41 |  | The permitted humidity limits are is  3% Too much low value. We propose 5% | Spain | Accepted |
| 42 |  | p.17 change „Om“ to „ohm“ (or its symbol) | Czech Republic | Accepted |
| 43 | p. 7.10.6 | Is it necessary to maintain the stationary sign during the 30 days or only it is applied to carry out the measures? | Spain | Not accepted It is necessary to maintain the stationary sign during 30 days. |
| 44 | p. 7.10.7 | To have equipments for all the ranges is very expensive and I don't believe that there are many pattern approvals. | Spain | Not accepted In our opinion the producer must point out for which ranges this test had been carried out and substantiate the assimilation of this results to other ranges. |
| 45 | p. 8.4.1 | The calibration is carried out by the direct comparison method - the comparison of characteristics of the calibrated instrument and the standard one, with the same pressure values and in the same ambient conditions not less than five testing cycles. instead of: five better : at least three cycles | Germany | Not accepted For determination of uncertainty not less than ten |
| 46 | p. 9.7.3 | “The verification Certificate for the transmitter shall, according to National Regulations, include the following:……………” In this list we miss the value for the expanded measurement uncertainty (U, for k=2) | Austria | Not accepted In this section only requirements to verification certificate are represented. During this procedure only intrinsic deviation (not uncertainty) is defined. |
| 47 | Annex C | Table “B.4” are constituents of uncertainty no mi | Spain | Accepted |
| 48 |  | Points 1 to 11 in Determination of uncertainty and its constituents is the same that in “Determination of intrinsic deviation and its constituents”. What is its contribution at the procedure? | Spain | Not accepted In ever concrete examination only one of two methods of determination accuracy parameters is used – either uncertainty or intrinsic deviation. But the constituents in both methods are the same. |
| 49 |  | change Cyrillic “i” into “and” | Czech Republic | Accepted |
| 50 |  | Formula 1.3, Formula 1.5, Formula 2.3, Formula 2.4, Formula 2.5 – Ypi is not defined | Austria | Accepted |
| 51 | General | Why is “with elastic sensing elements”? Do all the transducers have elastic sensing elements? It looks like as if only was used with bourdon, capsule…. Today, the most transmitters are with piezoelectric transducer. | Spain | Not accepted The transmitters without elastic sensing elements (for example – sensing elements made from gallium arsenide are known. |
| 52 |  | Suggest the third CD is numbered as proposed BIML templates and the IR is subdivided as Parts 1, 2 and 3. | United Kingdom | Accepted |
| 53 |  | The section on Metrological Requirements, 4.2. , seems too undefined permitting requirements to be drawn from manufacturer’s material and advertisements. Would it not be possible to define a number of accuracy classes and therefore define in a table the allowable deviations for each class? | United Kingdom | Not accepted This Recommendation is based on the documents wherein other methods of accuracy parameters determination are used. |
| 54 |  | The numbering is inconsequent in this second draft | Germany | Accepted |
| 55 |  | Throughout the document, mixed usage of decimal comma and decimal point (e.g. page 8)! | Czech Republic | Accepted |
| 56 | Bibliography | instead of: German Guidance DKD R 6-1 “Calibration of instruments for measuring pressure” the right title is: for DKD R 6-1 is “Calibration of Pressure Gauges” - ? Also?: - IEC 61298-1:1995-07, Process measurement and control devices – General methods and Procedures for evaluating performance – Part 1: General considerations. IEC 61298-2: 1995-07, Process measurement and control devices – General methods and Procedures for evaluating performance – Part 2: Tests under reference conditions. IEC 61298-3:1998-02, Process measurement and control devices – General methods and Procedures for evaluating performance – Part 3: Tests for the effects of influence quantities. | Germany | Accepted |