|  |  | COMMITTEE DRAFT OIML/CD1 Date: 19 April 2012 Reference number: OIML/TC12/R46-3/1 Supersedes document: - |
| --- | --- | --- |
| OIML TC 12 |  | Circulated to P- and O-members and liaison international bodies and external organizations for: |
| Title: Active Electrical Energy Meters |  | discussion at (date and place of meeting): ............. |
| Secretariat: Australia |  | comments by 19 July 2012: X |
|  |  | vote CIML Members by XXX |

TITLE OF THE CD (English):

OIML R 46-3

Active Electrical Energy Meters

Part 3: Test Report Format

TITRLE DU CD (French):

OIML R 46-3

Active Electrical Energy Meters

Part 3: Test Report Format

Original version in: English

- Contents

1	Information	4

1.1	Meter specification	4

1.2	Test values	4

2	General Requirements	5

2.1	Timing requirements (3.4)	6

3	Validation procedure (protection of metrological properties) (4.3; 3.6)	7

4	Tests for maximum permissible error	9

4.1	Initial intrinsic error for positive and negative flow (6.2.1)	9

4.2	Reverse energy flow (6.2.1)	10

4.3	Self heating (6.2.2)	11

4.4	Starting current (6.2.3)	12

4.5	Test of no-load condition (6.2.4)	12

4.6	Meter constants (6.2.5)	13

5	Tests for influence quantities	14

5.1	Temperature dependence (6.3.2; Table 3)	14

5.2	Load balance (6.3.3)	15

5.3	Voltage variation (6.3.4)	16

5.4	Frequency variations (6.3.5)	16

5.5	Harmonics in voltage and current (6.3.6)	17

5.6	Tilt (6.3.7)	17

5.7	Severe voltage variations (6.3.8)	18

5.8	One or two phases interrupted (6.3.9)	19

5.9	Sub-harmonics in the AC current circuit (6.3.10)	20

5.10	Harmonics in the AC current circuit (6.3.11)	20

5.11	Reversed phase sequence (any two phases interchanged) (6.3.12)	21

5.12	Continuous (DC) magnetic induction of external origin (6.3.13)	22

5.13	Magnetic field (AC, power frequency) of external origin (6.3.14)	23

5.14	Radiated, radio frequency (RF), electromagnetic fields (6.3.15.1)	24

5.15	Immunity to Conducted Disturbances, Induced by Radiofrequency Fields (6.3.15.2)	25

5.16	DC in the AC current circuit (6.3.16)	26

5.17	High-order harmonics  (6.3.17)	27

6	Test for disturbances	28

6.1	Critical change value (6.4.1 a); 3.3.6.2)	28

6.2	Magnetic field (AC, power frequency) of external origin (6.4.2)	28

6.3	Electrostatic discharge (6.4.3)	29

6.4	Fast transients (6.4.4)	30

6.5	Voltage dips and interruptions (6.4.5)	31

6.6	Radiated, radio frequency (RF), electromagnetic fields (6.4.6)	32

6.7	Surges on AC mains power lines (6.4.7)	33

6.8	Damped oscillatory waves immunity test (6.4.8)	34

6.9	Short-time overcurrent (6.4.9)	35

6.10	Impulse Voltage (6.4.10)	36

6.11	Earth Fault (6.4.11)	37

6.12	Operation of auxiliary devices (6.4.12)	38

6.13	Vibrations (6.4.13.1)	39

6.14	Shock (6.4.13.2)	40

6.15	Protection against solar radiation (6.4.14)	41

6.16	Protection against ingress of dust (6.4.15)	42

6.17	Extreme temperatures - Dry Heat (6.4.16.1)	43

6.18	Extreme temperatures - Cold (6.4.16.2)	44

6.19	Damp Heat, steady-state (non-condensing), for humidity class H1 (6.4.16.3)	45

6.20	Damp Heat, cyclic (condensing), for humidity class H2 and H3 (6.4.16.4)	46

6.21	Water test (6.4.16.5)	47

6.22	Durability (6.4.17)	48

# Information

## Meter specification

| Application no.: |  |
| --- | --- |
| Meter manufacturer: |  |
| Meter model: |  |
| Serial number: |  |
| Meter type: |  |

| Accuracy class: |  |  | A |  | B |  | C |  | D |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Nominal voltage, Unom: |  | V |
| --- | --- | --- |
| Nominal frequency, fnom: |  | Hz |
| Maximum current, Imax: |  | A |
| Transitional current, Itr: |  | A |
| Minimum current, Imin: |  | A |

|  |  | Direct-connected |  | Current transformer |  | Current and voltage transformers |
| --- | --- | --- | --- | --- | --- | --- |

| Connection mode (phases, wires, elements): |  |
| --- | --- |
| Alternative connection mode(s): |  |

Direction of energy flow / registers:

|  |  | Single-register, bi-directional |  | Single-register, positive direction only |
| --- | --- | --- | --- | --- |

|  |  | Two-register, bi-directional |  | Single-register, uni-directional |
| --- | --- | --- | --- | --- |

| Register multiplier: |  |  |
| --- | --- | --- |
| Meter constant: |  | (include units of measurement) |
| Specified clock frequencies: |  | (include units of measurement) |

| Lower specified temperature: |  |  | -55 °C |  | -40 °C |  | -25 °C |  | -10 °C |  | +5 °C |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Upper specified temperature: |  |  | +30 °C |  | +40 °C |  | +55 °C |  | +70 °C |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Humidity class: |  |  | H1 |  | H2 |  | H3 |
| --- | --- | --- | --- | --- | --- | --- | --- |

| Tilt / Mounting position: |  |  | Mounting position specified |  | Any position is allowed |
| --- | --- | --- | --- | --- | --- |

| Hardware version: |  |
| --- | --- |
| Software version: |  |

Remarks:

## Test values

When ranges of values are specified by the manufacturer, the values used for testing shall be specified below.

| Test voltage: |  | V |
| --- | --- | --- |
| Test frequency: |  | Hz |
| Test connection mode: |  |  |

Remarks:

# General Requirements

| Clause | Description | Passed | Failed | Remarks |
| --- | --- | --- | --- | --- |
| 3.1 | Units of measurement |  |  |  |
|  | Valid units of measurement (Wh, kWh, MWh, GWh) |  |  |  |
| 3.2; Table 1 | Rated operating conditions (Table 1) |  |  |  |
|  | Check Imax/Itr ratio complies |  |  |  |
|  | Check Imax/Imin ratio complies |  |  |  |
|  | Check Imax/Ist ratio complies |  |  |  |
| 3.6.9 | Checking facility event record |  |  |  |
|  | Check any checking facility for sufficient room for events and that it is of the first-in-first-out type. |  |  |  |
| 3.7.1 | Readability of result |  |  |  |
|  | Indicating device is easy to read |  |  |  |
|  | Height of characters of measurement result ≥ 4 mm |  |  |  |
|  | Any decimal fractions are clearly indicated |  |  |  |
|  | Able to display all data relevant for billing purposes |  |  |  |
|  | All registers relevant for billing can store and display energy = (4000·Unom·Imax·n) Wh, where n is the number of phases. (4000 h). |  |  |  |
|  | For mechanical registers |  |  |  |
|  | All decimal fraction drums are marked differently |  |  |  |
|  | For auto-sequencing displays |  |  |  |
|  | Each register for billing purposes is retained for ≥ 5 s |  |  |  |
|  | For multi-tariff meters |  |  |  |
|  | The register which reflects the active tariff is indicated |  |  |  |
|  | It is possible to read each tariff register locally and each register is clearly identified |  |  |  |
|  | For electronic registers |  |  |  |
|  | Retention time for results for a disconnected meter is ≥ 1 year |  |  |  |
|  | Electronic indicating devices are provided with a display test |  |  |  |
| 3.7.2 | Testability |  |  |  |
|  | The meter is equipped with a test output |  |  |  |
|  | The wavelength of radiated signals is between 550 nm to 1000 nm. |  |  |  |
|  | The radiation strength ET complies with limits at on and off conditions. |  |  |  |

## Timing requirements (3.4)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Limits shall be determined from IEC 62054-21 based on clock type.

| Test | Temperature (°C) | Duration | Result (s/day) | Limit (s/day) |
| --- | --- | --- | --- | --- |
| Mains operation | 23 | 30 days |  |  |
| High temperature: | 45 | 24 hours |  |  |
| Low temperature: | -10 | 24 hours |  |  |
| Operation reserve: | - | 36 hours |  |  |

- Check that each ¦result¦ ≤ ¦limit¦

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

# Validation procedure (protection of metrological properties) (4.3; 3.6)

| Application no.: |  | Observer: |  |
| --- | --- | --- | --- |
| Meter model: |  | Date: |  |

- The two applicable validation procedures are as follows: - AD: Analysis of the documentation and validation of the design. - VFTSw: Validation by functional testing of software functions.

| Clause | Requirements | Validation Description | Passed | Failed |
| --- | --- | --- | --- | --- |
| 3.6.2 | Software identification (AD + VFTSw) |  |  |  |
| Specify software identification and means of identification. |  |  |  |  |
| Validate the presentation or display of the software identification. |  |  |  |  |
| Validate that the software identification is inextricably linked to the software |  |  |  |  |
| 3.6.3.1 | Prevention misuse (AD + VFTSw) |  |  |  |
| Validate that possibilities of misuse are minimal. |  |  |  |  |
| 3.6.3.2 | Fraud protection (AD + VFTSw) |  |  |  |
| Validate that legally relevant software is secured against modification, loading or changes. |  |  |  |  |
| Validate that only clearly documented functions can be activated by the user interface. |  |  |  |  |
| Validate protection/sealing that makes unauthorised access impossible or evident. |  |  |  |  |
| 3.6.4 | Parameter protection (AD + VFTSw) |  |  |  |
| Validate that legally relevant characteristics are secured against unauthorised modification. |  |  |  |  |
| The following are considered as modifications to (legally relevant) device-specific parameters. - Zeroing or changing the register for total energy - Zeroing or changing the event record of a checking facility. |  |  |  |  |
| Validate that the meter stops registering energy when modifying a (legally relevant) device-specific parameter. |  |  |  |  |
| Validate (where applicable) a facility to record adjustments to device-specific parameters. |  |  |  |  |
| 3.6.5 | Separation of electronic devices and sub-assemblies (AD) |  |  |  |
| Identify the legally relevant part(s) of the meter. |  |  |  |  |
| Validate the separation. Metrologically critical parts of an electricity meter – whether software or hardware parts – shall not be inadmissibly influenced by other parts of the meter. |  |  |  |  |
| 3.6.6 | Separation of software (AD) |  |  |  |
| Identify and validate legally relevant software. |  |  |  |  |
| Identify and validate the interface between legally relevant software and other software parts. |  |  |  |  |
| Identify and validate documented interface commands and statement of completeness. |  |  |  |  |
| 3.6.7 | Storage of data, transmission via communication systems (AD + VFTSw) |  |  |  |
| Refer to clause 3.6.7 for applicability of these requirements. |  |  |  |  |
| Validate that measurement values stored or transmitted are accompanied by all information necessary for future legally relevant use. |  |  |  |  |
| 3.6.7.1.2 | Data protection with respect to time of measurement (AD + VFTSw) |  |  |  |
| Validate that software data protection with respect to time of measurement. |  |  |  |  |
| 3.6.7.2 | Automatic storing (AD + VFTSw) |  |  |  |
| Validate automatic storage of data. |  |  |  |  |
| Validate sufficient permanency and memory for storage of data. |  |  |  |  |
| Validate deletion of stored data. |  |  |  |  |
| 3.6.7.3.1 | Transmission delay (AD + VFTSw) |  |  |  |
| Validate that measurement is not inadmissibly influenced by a transmission delay. |  |  |  |  |
| 3.6.7.3.2 | Transmission interruption (AD + VFTSw) |  |  |  |
| Validate measurement data is not lost due to unavailability of network services. |  |  |  |  |
| 3.6.7.4 | Time stamp (AD + VFTSw) |  |  |  |
| Validate that time stamps are read from the clock of the device. |  |  |  |  |
| Validate that setting of the clock is protected as a legally relevant parameter. |  |  |  |  |
| 3.6.8 | Maintenance and re-configuration (AD + VFTSw) |  |  |  |
| Identify and validate the implementation for software updates. |  |  |  |  |

# Tests for maximum permissible error

## Initial intrinsic error for positive and negative flow (6.2.1)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- If a meter is specified with alternate connection modes, this test shall be made for all specified connection modes.

| Connection mode: |  |
| --- | --- |

| IX: testpoint specified by the national authority between Itr and Imax: |  |
| --- | --- |

| Positive energy flow |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Test Current (A) | Power Factor | Error (%) |  | Mean error (%) | Base m.p.e. (%) |
|  |  | ↓ | ↑ |  |  |
| Imin | unity |  |  |  |  |
| Itr |  |  |  |  |  |
| IX |  |  |  |  |  |
| Imax |  |  |  |  |  |
| Imin | (most inductive) |  |  |  |  |
| Itr |  |  |  |  |  |
| IX |  |  |  |  |  |
| Imax |  |  |  |  |  |
| Imin | (most capacitive) |  |  |  |  |
| Itr |  |  |  |  |  |
| IX |  |  |  |  |  |
| Imax |  |  |  |  |  |
| Negative energy flow |  |  |  |  |  |
| Itr | unity |  |  |  |  |
| Imax |  |  |  |  |  |
| Itr | (most inductive) |  |  |  |  |
| Imax |  |  |  |  |  |
| Itr | (most capacitive) |  |  |  |  |
| Imax |  |  |  |  |  |

- Check that each ¦mean error¦ ≤ ¦base m.p.e¦

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Reverse energy flow (6.2.1)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

| Calculation for test time for reverse flow |  | Imin | Imax |
| --- | --- | --- | --- |
| a) | Time that the test output would register ten pulses in the forward energy flow direction (minutes): |  |  |
| b) | Time that the primary register would register 2 units of the least significant digit in the forward energy flow direction (minutes): |  |  |
| c) | 1 minute: | 1 | 1 |
| Test time is the maximum of a), b) and c): |  |  |  |

| Test Current (A) | Power Factor | Test Time (minutes) | Change in register |  | Number of test pulses |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Measured | Limit | Measured | Limit |
| Imin | unity |  |  | 0 |  | 1 |
| Imax |  |  |  |  |  |  |

- Check that there is no change in the energy registered in the primary register. - Check that the number of test pulses emitted ≤ 1.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Self heating (6.2.2)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

| Voltage circuits energised for time: |  | (At least 1 hour for Class A, 2 hours for all other classes.) |
| --- | --- | --- |

- The test shall be carried out for at least 1 hour, and in any event until the variation of error over any 20-minute period does not exceed 10 % of base maximum permissible error.

| Test current (A) | Power factor | Time at Imax (minutes) | Error (%) | Base m.p.e. (%) | Error shift (%) | Limit (%) (½ of base m.p.e.) |
| --- | --- | --- | --- | --- | --- | --- |
| Imax | Unity |  |  |  |  |  |

| Has the error shift levelled out? |  | If no, continue test according to (a) or (b) below. |
| --- | --- | --- |

(a) If the load can be changed in less than 30 seconds, then:

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) | Error shift (%) | Limit (%) (½ of base m.p.e.) |
| --- | --- | --- | --- | --- | --- |
| Imax | 0.5 inductive |  |  |  |  |

(b) Else, allow meter to return to its initial temperature and repeat test for power factor 0.5 inductive.

| Voltage circuits energised for time: |  | (At least 1 hour for Class A, 2 hours for all other classes.) |
| --- | --- | --- |

| Test current (A) | Power factor | Time at Imax (minutes) | Error (%) | Base m.p.e. (%) | Error shift (%) | Limit (%) (½ of base m.p.e.) |
| --- | --- | --- | --- | --- | --- | --- |
| Imax | 0.5 inductive |  |  |  |  |  |

- Check that each ¦error¦ ≤ ¦base m.p.e¦ - Check that each ¦error shift¦ ≤ ¦limit¦

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Starting current (6.2.3)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Determine the error at the starting current based on the rate of test pulses (or revolutions if no test output).

| Expected time between pulses, τ=3.6⋅k/(m⋅Unom⋅Ist) seconds: |  |
| --- | --- |

| Test current (A) | Power factor | Meter started (Yes/No) | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- | --- |
|  | Unity |  |  |  |

- Check that the ¦error¦ ≤ ¦base m.p.e¦

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Test of no-load condition (6.2.4)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

| Minimum test period, Δt≥100×103/(b⋅k⋅m⋅Unom⋅Imin) hours: |  |
| --- | --- |

| Test current (A) | Test period Δt (hours) | For meters with a test output |  | For electromechanical meters |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Number of pulses emitted | Limit | Rotor revolutions | Limit |
| No current |  |  | 1 |  | Less than a complete revolution |

- For meters with a test output, check if the number of pulses emitted ≤ 1. - For electromechanical meters, check that the rotor does not make a complete revolution.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Meter constants (6.2.5)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

| Does the meter have multiple registers or pulse outputs under legal control? (Yes/No) |  |
| --- | --- |
| If yes, is there a system in place to guarantee identical behaviour of meter constants? (Yes/No) |  |
| If yes, specify the system, otherwise all registers and pulse outputs must be tested. |  |

| Register and test output tested: |  |
| --- | --- |
| Apparent resolution of basic energy register, R: |  |
| Minimum energy to be passed through, Emin=1000⋅R/b: |  |

| Test current (A) | Power factor | Energy measured by |  | Relative difference (%) t−r/r | Limit (%) (10% of base m.p.e.) |
| --- | --- | --- | --- | --- | --- |
|  |  | Register (r) | Test output (t) |  |  |
|  | Unity |  |  |  |  |

- Check that each ¦relative difference¦ ≤ ¦limit¦

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

# Tests for influence quantities

## Temperature dependence (6.3.2; Table 3)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- The mean temperature coefficient, c, is calculated by c=(eu−el)/(tu−tl). - Temperature intervals shall span at least 15°C and no more than 23°C. - The set of intervals must span the entire specified operating range (intervals may overlap). - A temperature coefficients table must be completed for each temperature interval.

| Temperature coefficients table |  |  | Temperature interval (tl to tu): |  | tl (°C): |  |  | tu (°C): |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test Current (A) | Power factor | Error (%) |  |  |  | Mean temperature coefficient (%/K) |  |  |  |  |
|  |  | el |  | eu |  | c |  |  | Limit |  |
| Itr | unity |  |  |  |  |  |  |  |  |  |
| 10 Itr |  |  |  |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |  |  |  |
| Itr | 0.5 capacitive |  |  |  |  |  |  |  |  |  |
| 10 Itr |  |  |  |  |  |  |  |  |  |  |
| Imax |  |  |  |  |  |  |  |  |  |  |

- Check that each ¦c¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Load balance (6.3.3)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

This test is only for poly-phase meters and for single-phase three-wire meters.

- Reference voltages shall be applied to all voltage circuits

| Test current (A) | Power factor | Load | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Balanced |  |  |  |
|  |  | Current in L1 only |  |  |  |
|  |  | Current in L2 only |  |  |  |
|  |  | Current in L3 only |  |  |  |
| Imax | unity | Balanced |  |  |  |
|  |  | Current in L1 only |  |  |  |
|  |  | Current in L2 only |  |  |  |
|  |  | Current in L3 only |  |  |  |
| 10 Itr | 0.5 inductive | Balanced |  |  |  |
|  |  | Current in L1 only |  |  |  |
|  |  | Current in L2 only |  |  |  |
|  |  | Current in L3 only |  |  |  |
| Imax | 0.5 inductive | Balanced |  |  |  |
|  |  | Current in L1 only |  |  |  |
|  |  | Current in L2 only |  |  |  |
|  |  | Current in L3 only |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Voltage variation (6.3.4)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- If several Unom values are stated, the test shall be repeated for each Unom value.

| Unom (V): |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) |  | Power factor |  | Voltage variation | Error (%) | Error shift (%) | Limit (%) |
| 10 Itr |  | unity |  | Reference (Unom) |  |  |  |
|  |  |  |  | 0.9 Unom |  |  |  |
|  |  |  |  | 1.1 Unom |  |  |  |
| 10 Itr |  | 0.5 inductive |  | Reference (Unom) |  |  |  |
|  |  |  |  | 0.9 Unom |  |  |  |
|  |  |  |  | 1.1 Unom |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Frequency variations (6.3.5)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- If several fnom values are stated, the test shall be repeated for each fnom value.

| fnom (Hz): |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) |  | Power factor |  | Frequency variation | Error (%) | Error shift (%) | Limit (%) |
| 10 Itr |  | unity |  | Reference (fnom) |  |  |  |
|  |  |  |  | 0.98 fnom |  |  |  |
|  |  |  |  | 1.02 fnom |  |  |  |
| 10 Itr |  | 0.5 inductive |  | Reference (fnom) |  |  |  |
|  |  |  |  | 0.98 fnom |  |  |  |
|  |  |  |  | 1.02 fnom |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Harmonics in voltage and current (6.3.6)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Determine the error shift, relative to the error at reference conditions (with no harmonics), when the quadriform waveform (Table 11), is applied to both voltage and current circuits. - Determine the error shift, relative to the error at reference conditions (with no harmonics), when the peaked waveform (Table 12), is applied to both voltage and current circuits.

| Test current (A) | Power factor | Harmonics applied to both voltage and current circuits | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (fnom) |  |  |  |
|  |  | Quadriform waveform |  |  |  |
|  |  | Peaked waveform |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Tilt (6.3.7)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

This test is only for electromechanical meters or meters of other constructions that may be influenced by the working position.

| Operating position specified by manufacturer: |  |
| --- | --- |
| Define perpendicular orientations corresponding to forward, backward, left and right. |  |

| Test current (A) | Power factor | Tilt | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| Itr | unity | Reference (no tilt) |  |  |  |
|  |  | 3° forward |  |  |  |
|  |  | 3° backward |  |  |  |
|  |  | 3° left |  |  |  |
|  |  | 3° right |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Severe voltage variations (6.3.8)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- If several Unom values are stated, the test shall be repeated for each Unom value.

Test procedure 1

| Unom (V): |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) |  | Power factor |  | Voltage variation | Error (%) | Error shift (%) | Limit (%) |
| 10 Itr |  | unity |  | Reference (Unom) |  |  |  |
|  |  |  |  | 0.8 Unom |  |  |  |
|  |  |  |  | 0.85 Unom |  |  |  |
|  |  |  |  | 1.15 Unom |  |  |  |

Test procedure 2

| Does the meter have distinct shut-down / turn-on voltages? (Yes/No) |  |
| --- | --- |
| Shut-down voltage (V): |  |
| Turn-on voltage (V): |  |
| If yes, two additional mandatory testpoints (shutdown low and shutdown high) shall be included. Shutdown low shall be within a 2 V range below the shut-down voltage. Shutdown high shall be within a 2 V range above the turn-on voltage. |  |

| Unom (V): |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) |  | Power factor |  | Voltage variation | Error (%) | Error shift (%) | Limit (%) |
| 10 Itr |  | unity |  | Reference (Unom) |  |  |  |
|  |  |  |  | 0.7 Unom |  |  | +10 to -100 |
|  |  |  |  | 0.6 Unom |  |  |  |
|  |  |  |  | 0.5 Unom |  |  |  |
|  |  |  |  | 0.4 Unom |  |  |  |
|  |  |  |  | 0.3 Unom |  |  |  |
|  |  |  |  | 0.2 Unom |  |  |  |
|  |  |  |  | 0.1 Unom |  |  |  |
|  |  |  |  | 0 Unom |  |  |  |
|  |  |  |  | shutdown low |  |  |  |
|  |  |  |  | shutdown high |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## One or two phases interrupted (6.3.9)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

This test is only for poly-phase meters with three measuring elements

- One or two phases are removed while keeping the load current constant.

| Test current (A) | Power factor | Load | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (no phases removed) |  |  |  |
|  |  | Phase L1 removed |  |  |  |
|  |  | Phase L2 removed |  |  |  |
|  |  | Phase L3 removed |  |  |  |
|  |  | Phases L1, L2 removed |  |  |  |
|  |  | Phases L1, L3 removed |  |  |  |
|  |  | Phases L2, L3 removed |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Sub-harmonics in the AC current circuit (6.3.10)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- The sub-harmonic waveform is formed from a sinusoidal signal with twice the amplitude of the reference signal, which is switched on and off every second period (as shown in Figure 3 b)).

| Test current (A) | Power factor | Current signal | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (sinusoidal, fnom) |  |  |  |
|  |  | Sub-harmonic waveform |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Harmonics in the AC current circuit (6.3.11)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- The harmonic waveform is formed from a sinusoidal signal with twice the amplitude of the reference signal, which is set to zero during the first and third quarters of the period.

| Test current (A) | Power factor | Current signal | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (sinusoidal, fnom) |  |  |  |
|  |  | Harmonic waveform |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Reversed phase sequence (any two phases interchanged) (6.3.12)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- This test only applies to three-phase meters.

| Test current (A) | Power factor | Phase sequence | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (L1, L2, L3) |  |  |  |
|  |  | L1, L3, L2 |  |  |  |
|  |  | L2, L1, L3 |  |  |  |
|  |  | L3, L2, L1 |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Continuous (DC) magnetic induction of external origin (6.3.13)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Permanent magnet with a surface area of at least 2000 mm2. - Field along axis of magnet’s core at 30 mm from surface: 200 mT ± 30 mT. - 6 points per meter surface. Report greatest error shift for each surface.

Specify or illustrate the surfaces designated as front, back, top, bottom, left and right.

| Test current (A) | Power factor | Meter surface tested | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (no magnetic induction) |  |  |  |
|  |  | Front |  |  |  |
|  |  | Back |  |  |  |
|  |  | Top |  |  |  |
|  |  | Bottom |  |  |  |
|  |  | Left |  |  |  |
|  |  | Right |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Magnetic field (AC, power frequency) of external origin (6.3.14)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Continuous field, 400 A/m, f = fnom - Field at three orthogonal directions.

Specify or illustrate the three orthogonal directions relative to the meter designated as x, y & z:

| Test current (A) | Power factor | Magnetic field axis direction | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Reference (no magnetic induction) |  |  |  |
|  |  | x-axis |  |  |  |
|  |  | y-axis |  |  |  |
|  |  | z-axis |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Radiated, radio frequency (RF), electromagnetic fields (6.3.15.1)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to radiated radiofrequency fields.

Test Condition 1 – with current

- Frequency range: 80 to 6000 MHz - Field strength: 10 V/m - Modulation: 80% AM, 1 kHz sine wave - The meter shall be separately tested at the manufacturer’s specified clock frequencies. - Any other sensitive frequencies shall also be analysed separately.

| Test current (A) | Power factor | Antenna | Frequency value / range (MHz) | Polariz- ation | Facing meter | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 10 Itr | unity |  |  | Vertical | Front |  |  |
|  |  |  |  |  | Back |  |  |
|  |  |  |  |  | Right |  |  |
|  |  |  |  |  | Left |  |  |
|  |  |  |  |  | Top |  |  |
|  |  |  |  |  | Bottom |  |  |
|  |  |  |  | Horizontal | Front |  |  |
|  |  |  |  |  | Back |  |  |
|  |  |  |  |  | Right |  |  |
|  |  |  |  |  | Left |  |  |
|  |  |  |  |  | Top |  |  |
|  |  |  |  |  | Bottom |  |  |
|  |  | [extend for each antenna] |  |  |  |  |  |
|  |  |  | [extend for clock frequencies and any other sensitive frequencies] |  |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Immunity to Conducted Disturbances, Induced by Radiofrequency Fields (6.3.15.2)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to conducted disturbances induced by RF fields.

- Frequency range: 0.15 to 80 MHz - Field strength: 10 V (e.m.f.) - Modulation: 80% AM, 1 kHz sine wave - Test all power ports and I/O ports.

| Test current (A) | Power factor | Power or I/O port | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- |
| 10 Itr | unity |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## DC in the AC current circuit (6.3.16)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Electromechanical and transformer operated meters shall be assumed to be immune to DC in the AC current circuit.

| Test current (A) | Power factor | Current test wave | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- |
| Imax/2√2 | unity | Sinusoidal (intrinsic error) |  |  |  |
| Imax/√2 |  | Half-wave rectified |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## High-order harmonics	 (6.3.17)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Asynchronous test signals, swept from f = 15 fnom to 50 fnom. - Sweep from low frequency to high frequency, and then back down. - One reading shall be taken per harmonic frequency (report maximum error within the frequency range). - Report greatest error shift for each sweep.

Voltage Circuit Test

- Asynchronous test signal: 0.02 Unom

| Test current (A) | Power factor | Signal on voltage circuit | Sweep direction | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- | --- |
| Itr | unity | Sinusoidal (intrinsic error) |  |  |  |  |
|  |  | Test signal superimposed | low to high |  |  |  |
|  |  |  | high to low |  |  |  |

Current Circuit Test

- Asynchronous test signal: 0.1 Itr

| Test current (A) | Power factor | Signal on current circuit | Sweep | Error (%) | Error shift (%) | Limit (%) |
| --- | --- | --- | --- | --- | --- | --- |
| Itr | unity | Sinusoidal (intrinsic error) |  |  |  |  |
|  |  | Test signal superimposed | low to high |  |  |  |
|  |  |  | high to low |  |  |  |

- Check that each ¦error shift¦ ≤ ¦limit¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

# Test for disturbances

## Critical change value (6.4.1 a); 3.3.6.2)

| - The critical change value is used as a criterion for significant fault in many disturbance tests. |  |  |  |
| --- | --- | --- | --- |
| Number of measuring elements, m: |  |  |  |
| Nominal voltage, Unom: |  | V |  |
| Maximum current, Imax: |  | A |  |
| Critical change value (m⋅Unom⋅Imax⋅10−6): |  | kWh |  |

## Magnetic field (AC, power frequency) of external origin (6.4.2)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

- Magnetic field strength short duration (3 s): 1000 A/m, f = fnom - Voltage circuits energised with Unom. - No current in the current circuits. - Field at three orthogonal directions.

Specify the three orthogonal directions relative to the meter designated as x, y & z:

a) Check for significant fault (see critical change value in 6.1)

| Magnetic field axis direction | Change in… |  | Critical change value |
| --- | --- | --- | --- |
|  | Register | Equivalent energy of the test output |  |
| x-axis |  |  |  |
| y-axis |  |  |  |
| z-axis |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Electrostatic discharge (6.4.3)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to electrostatic discharges.

- Contact discharge is the preferred test method. Air discharges shall be used where contact discharge cannot be applied. - Voltage circuits energised with Unom. - Current and auxiliary circuits open, with no current.

a) Check for significant fault (see critical change value in 6.1)

| Application | Discharge mode | Test voltage (kV) | Polarity | Number of discharges (≥ 10) | Change in… |  | Critical change value |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Register | Equivalent energy of the test output |  |
| Direct | Contact | 8 | Positive |  |  |  |  |
|  |  |  | Negative |  |  |  |  |
|  | Air | 15 | Positive |  |  |  |  |
|  |  |  | Negative |  |  |  |  |
| Indirect, Horizontal coupling plane | Contact | 8 | Positive |  |  |  |  |
|  |  |  | Negative |  |  |  |  |
| Indirect, Vertical coupling plane | Contact | 8 | Positive |  |  |  |  |
|  |  |  | Negative |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Fast transients (6.4.4)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to fast transients.

- The test voltage shall be applied in common mode (line-to-earth) to: a) the voltage circuits; b) the current circuits, if separated from the voltage circuits in normal operation; c) the auxiliary circuits, if separated from the voltage circuits in normal operation and with a reference voltage over 40 V.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| Test current (A) | Power factor | Circuit / Auxiliary circuit | Test Voltage (kV) | Error (%) | Error shift (%) | Limit of error shift (%) |
| --- | --- | --- | --- | --- | --- | --- |
| 10 Itr | unity | Voltage | 4 |  |  |  |
|  |  | Current |  |  |  |  |
|  |  | [Auxiliary circuits] | 2 |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Voltage dips and interruptions (6.4.5)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to voltage dips and interruptions.

- Voltage circuits energised with Unom. - Without current in the current circuit.

a) Check for significant fault (see critical change value in 6.1)

| Dip / Interruption | Test | Amplitude relative to Unom | Duration (cycles) | Repetitions | Change in… |  | Critical change value |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Register | Equivalent energy of the test output |  |
| Dip | Test a | 30% | 0.5 | 10 |  |  |  |
|  | Test b | 60% | 1 | 10 |  |  |  |
|  | Test c | 60% | [25/30] [1] | 10 |  |  |  |
| Interruption | - | 0% | [250/300] [2] | - |  |  |  |

Note [1]: Duration (cycles) for Voltage Dip Test c depends on the reference frequency 25 for 50 Hz, 30 for 60 Hz.

Note [2]: Duration (cycles) for Voltage Interruption Test depends on the reference frequency 250 for 50 Hz, 300 for 60 Hz.

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Radiated, radio frequency (RF), electromagnetic fields (6.4.6)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

Meters, such as electromechanical meters, which have been constructed using only passive elements shall be assumed to be immune to radiated radiofrequency fields.

Test Condition 2 – without current

- Voltage circuits energised with Unom, auxiliary circuits energized with reference voltage. - Without current in the current circuits and with the current circuits open-circuited. - Otherwise conditions as specified for the influence test with current in 5.14 above.

a) Check for significant fault (see critical change value in 6.1)

| Antenna | Frequency value / range (MHz) | Polariz- ation | Facing meter | Change in… |  | Critical change value |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Register | Equivalent energy of the test output |  |
|  |  | Vertical | Front |  |  |  |
|  |  |  | Back |  |  |  |
|  |  |  | Right |  |  |  |
|  |  |  | Left |  |  |  |
|  |  |  | Top |  |  |  |
|  |  |  | Bottom |  |  |  |
|  |  | Horizontal | Front |  |  |  |
|  |  |  | Back |  |  |  |
|  |  |  | Right |  |  |  |
|  |  |  | Left |  |  |  |
|  |  |  | Top |  |  |  |
|  |  |  | Bottom |  |  |  |
| [extend for each antenna] |  |  |  |  |  |  |
|  | [extend for clock frequencies and any other sensitive frequencies] |  |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Surges on AC mains power lines (6.4.7)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

This test is not applicable for meters such as electromechanical meters which shall be assumed to be immune to surges.

- Without any current in the current circuits and the current terminals open. - Number of tests: 5 positive and 5 negative - Repetition rate: maximum 1 per minute

a) Check for significant fault (see critical change value in 6.1)

| Amplitude (kV) | Application |  | Angle | Polarity | Change in… |  | Critical change value |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Register | Equivalent energy of the test output |  |
| Voltage Circuits |  |  |  |  |  |  |  |
| 2 | Line to line |  | 60° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
|  |  |  | 240° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
| 4 | Line to earth (1) |  | 60° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
|  |  |  | 240° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
| Auxiliary Circuits with a reference voltage over 40V (Repeat table below for each auxiliary circuit) |  |  |  |  |  |  |  |
| Specify auxiliary Circuit: |  |  |  |  |  |  |  |
| 1 | Line to line |  | 60° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
|  |  |  | 240° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
| 2 | Line to earth (1) |  | 60° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |
|  |  |  | 240° | Positive |  |  |  |
|  |  |  |  | Negative |  |  |  |

(1) For cases where the earth of the meter is separate to neutral.

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Damped oscillatory waves immunity test (6.4.8)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time (hh:mm): |  |  |

This test is only for meters intended to be operated with voltage transformers.

- Test duration: 60 s (15 cycles with 2 s on, 2 s off, for each frequency)

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Mode |  | Test Voltage (kV) | Test frequency (kHz) | Repetition rate (Hz) | Initial Error (%) | Error (%) | Error shift (%) | Limit of error shift (%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Voltage Circuits |  |  |  |  |  |  |  |  |  |  |
| 20 Itr | unity | Common |  | 2.5 | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | 0.5 inductive |  |  |  | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | unity | Differential |  | 1.0 | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | 0.5 inductive |  |  |  | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
| Auxiliary Circuits with a reference voltage over 40V (Repeat table below for each auxiliary circuit) |  |  |  |  |  |  |  |  |  |  |
| Specify auxiliary Circuit: |  |  |  |  |  |  |  |  |  |  |
| 20 Itr | unity | Common |  | 2.5 | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | 0.5 inductive |  |  |  | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | unity | Differential |  | 1.0 | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |
|  | 0.5 inductive |  |  |  | 100 | 40 |  |  |  |  |
|  |  |  |  |  | 1000 | 400 |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Short-time overcurrent (6.4.9)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- For direct connected meters: 30 Imax +0%, -10% for one half cycle at rated frequency. - For meters connected through current transformers: a current equivalent to 20 Imax +0%, -10%, for 0.5 s.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Phase | Intrinsic Error (%) |
| --- | --- | --- | --- |
| 10 Itr | unity | L1 |  |
|  |  | L2 |  |
|  |  | L3 |  |

| Application of overcurrent |  |  |  |  |  | After return to normal temperature |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Phase | Short-time overcurrent (A) | Duration | Damage caused? | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity | L1 |  |  |  |  |  |  |
|  |  | L2 |  |  |  |  |  |  |
|  |  | L3 |  |  |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Impulse Voltage (6.4.10)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- For each test, the impulse voltage is applied 10 times for each polarity. Minimum of 30 s between impulses.

a) Check for significant fault (see critical change value in 6.1)

| Test | Impulse Voltage (V) | Flashover, disruptive discharge or puncture? | Change in… |  | Critical change value |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Register | Equivalent energy of the test output |  |
| For circuits and between circuits |  |  |  |  |  |
| Circuits relative to earth |  |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that during the test, there is no flashover, disruptive discharge or puncture. - Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Earth Fault (6.4.11)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

This test only applies to three-phase four-wire transformer-operated meters connected to distribution networks which are equipped with earth fault neutralizers or in which the star point is isolated.

- Simulated earth fault condition in one of the three lines. - All voltages increased to 1.1 Unom. - Duration: 4 hours.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| Earth-fault condition |  |  |  |  | After return to normal temperature |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Voltage (V) | Duration (hours) | Damage caused? | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity | 1.1 Unom | 4 |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that after the test, the meter shows no damage. - Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Operation of auxiliary devices (6.4.12)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- Error continuously monitored while auxiliary devices are operated.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| Itr | unity |  |
| Imax |  |  |

| Test current (A) | Power factor | Auxiliary Device | Error (%) | Error shift (%) | Limit of error shift (%) |
| --- | --- | --- | --- | --- | --- |
| Itr | unity |  |  |  |  |
| Imax |  |  |  |  |  |
| Itr |  |  |  |  |  |
| Imax |  |  |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Vibrations (6.4.13.1)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- Meter mounted as in normal operation. - Vibrations applied, in turn, in three mutually perpendicular axes.

Specify or illustrate the three orthogonal directions relative to the meter designated as x, y & z:

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| After vibrations applied |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Shock (6.4.13.2)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- Meter not operational during tests.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| After shocks applied |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Protection against solar radiation (6.4.14)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

For outdoor meters only.

- Meter condition: non-operational during test. - Partially mask a section of the meter for later comparison. - Meter exposed to artificial radiation according to clause 6.4.14.

| Visual inspection requirements after exposure |  |  |
| --- | --- | --- |
| Clause | Check for effects on… | Remarks |
| (3.5) Markings on the meter | Legibility and permanency of markings |  |
| (3.6.1) Protection of metrological properties | Seals |  |
| (3.7.1) Readability of result | Transparent surfaces on indicating device |  |
|  | Indicating device |  |
| (3.3.6.2; Table 5) No alteration in appearance | Appearance |  |

a) Check for significant fault (see critical change value in 6.1)

| Change in… | Register |  |
| --- | --- | --- |
|  | Equivalent energy of the test output |  |
| Critical change value |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each visual inspection requirement is satisfied. - Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Protection against ingress of dust (6.4.15)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

| Visual inspection requirements after dust test |  |
| --- | --- |
| Visually inspect interior of meter | Remarks |
| Check if the talcum powder or other dust used in the test has accumulated in a quantity or location such that it could interfere with the correct operation of the equipment or impair safety. |  |
| Check that no dust has deposited where it could lead to tracking along the creepage distances. |  |

a) Check for significant fault (see critical change value in 6.1)

| Change in… | Register |  |
| --- | --- | --- |
|  | Equivalent energy of the test output |  |
| Critical change value |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each visual inspection requirement is satisfied. - Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Extreme temperatures - Dry Heat (6.4.16.1)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- Meter condition: non-operational.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| Dry Heat Test |  |
| --- | --- |
| Test temperature (one step higher than upper specified temperature) (°C) |  |
| Duration (hours) | 2 |

| After dry heat test |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Extreme temperatures - Cold (6.4.16.2)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

- Meter condition: non-operational.

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | unity |  |

| Cold Test |  |
| --- | --- |
| Test temperature (one step lower than lower specified temperature) (°C) |  |
| Duration (hours) | 2 |

| After cold test |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Damp Heat, steady-state (non-condensing), for humidity class H1 (6.4.16.3)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

For humidity class H1 only.

- Voltage and auxiliary circuits energised with reference voltage. - Without any current in the current circuits.

| Damp Heat, steady-state test |  |
| --- | --- |
| Temperature | 30 °C |
| Humidity | 85 % |
| Duration | 2 days |

a) Check for significant fault (limit of error shift and see critical change value in 6.1)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| 10 Itr | Unity |  |

| Change in… | Register |  |
| --- | --- | --- |
|  | Equivalent energy of the test output |  |
| Critical change value |  |  |

| Immediately after the test, check error shift according to Table 4 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks – 24 hours after the test

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e. – 24 hours after the test

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

Checks for damage or corrosion – 24 hours after test

| Requirement | Remarks |
| --- | --- |
| Check for evidence of any mechanical damage or corrosion which may affect the functional properties of the meter |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check that ¦error shift¦ ≤ ¦limit of error shift¦ immediately after the test. - Check all operational checks pass 24 hours after the test. - Check that ¦error¦ ≤ ¦base m.p.e.¦ 24 hours after the test. - Check that the requirements for damage or corrosion are satisified.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Damp Heat, cyclic (condensing), for humidity class H2 and H3 (6.4.16.4)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

For humidity class H2 or H3 only.

- Voltage and auxiliary circuits energised with reference voltage. - Without any current in the current circuits.

| Damp Heat, cyclic test |  |
| --- | --- |
| Specified humidity class |  |
| Lower temperature (°C) | 25 °C |
| Upper temperature (°C) |  |
| Duration | 2 cycles |

a) Check for significant fault (limit of error shift and see critical change value in 6.1)

| Test current (A) | Power factor | Initial Error (%) |
| --- | --- | --- |
| 10 Itr | Unity |  |

| Change in… | Register |  |
| --- | --- | --- |
|  | Equivalent energy of the test output |  |
| Critical change value |  |  |

| Immediately after the test, check error shift according to Table 4 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| 10 Itr | unity |  |  |  |

b) & c) Operational checks – 24 hours after the test

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e. – 24 hours after the test

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

Checks for damage or corrosion – 24 hours after test

| Requirement | Remarks |
| --- | --- |
| Check for evidence of any mechanical damage or corrosion which may affect the functional properties of the meter |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check that ¦error shift¦ ≤ ¦limit of error shift¦ immediately after the test. - Check all operational checks pass 24 hours after the test. - Check that ¦error¦ ≤ ¦base m.p.e.¦ 24 hours after the test. - Check that the requirements for damage or corrosion are satisified.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Water test (6.4.16.5)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

For humidity class H3 only.

- The meter shall be in functional mode, with no current.

| Water test |  |
| --- | --- |
| Flow rate (per nozzle): | 0.07 L/min |
| Angle of inclination: | 0° and 180° |
| Duration | 10 minutes |

a) Check for significant fault (see critical change value in 6.1)

| Change in… | Register |  |
| --- | --- | --- |
|  | Equivalent energy of the test output |  |
| Critical change value |  |  |

Accuracy immediately after the test

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

b) & c) Operational checks – 24 hours after the test

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e. – 24 hours after the test

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

Checks for damage or corrosion – 24 hours after test

| Requirement | Remarks |
| --- | --- |
| Check for evidence of any mechanical damage or corrosion which may affect the functional properties of the meter |  |

- Check that each ¦change in register¦ ≤ critical change value. - Check that each ¦change in equivalent energy of the test output¦ ≤ critical change value. - Check that ¦error¦ ≤ ¦base m.p.e.¦ immediately after the test. - Check all operational checks pass 24 hours after the test. - Check that ¦error¦ ≤ ¦base m.p.e.¦ 24 hours after the test. - Check that the requirements for damage or corrosion are satisified.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks:

## Durability (6.4.17)

| Application no.: |  |  |  | At start | At end |
| --- | --- | --- | --- | --- | --- |
| Meter model: |  |  | Temperature (°C): |  |  |
| Accuracy class: |  |  | Relative Humidity (%): |  |  |
| Observer: |  |  | Atmos. Pressure (kPa): |  |  |
| Date: |  |  | Time: |  |  |

| Specify Durability Standard applied: |  |
| --- | --- |

Specify details of durability test:

a) Check for significant fault (limit of error shift)

| Test current (A) | Power factor | Intrinsic Error (%) |
| --- | --- | --- |
| Itr | unity |  |
| 10 Itr | unity |  |
| Imax | unity |  |

| After durability |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test current (A) | Power factor | Error (%) | Error shift (%) | Limit of error shift (%) |
| Itr | unity |  |  |  |
| 10 Itr | unity |  |  |  |
| Imax | unity |  |  |  |

b) & c) Operational checks

| Test current (A) | Power factor | b) Operational check… | c) Check correct operation of… |  |
| --- | --- | --- | --- | --- |
|  |  | Does meter register energy? | Pulse outputs? | Tariff change inputs? |
|  | unity |  |  |  |

d) Check base m.p.e.

| Test current (A) | Power factor | Error (%) | Base m.p.e. (%) |
| --- | --- | --- | --- |
| Itr | unity |  |  |
| 10 Itr | 0.5 inductive |  |  |

- Check that each ¦error shift¦ ≤ ¦limit of error shift¦. - Check all operational checks pass. - Check that ¦error¦ ≤ ¦base m.p.e.¦.

|  |  | Passed |  | Failed |
| --- | --- | --- | --- | --- |

Remarks: